## Slide 1: FNAL Testbeam Results on LYSO+SiPM
CMS Timing Meeting
05/26/2016
California Institute of Technology
Adi Bornheim, Jay Lawhorn, Aashrita Mangu, Cristián Peña, Maria Spiropulu, Si Xie, Zhicai, Zhang
FNAL
Artur Apresyan, Sergey Los

## Slide 2: FNAL Testbeam Results on LYSO+SiPM
CMS Timing Meeting
05/26/2016
California Institute of Technology
Adi Bornheim, Jay Lawhorn, Aashrita Mangu, Cristián Peña, Maria Spiropulu, Si Xie, Zhicai, Zhang
FNAL
Artur Apresyan, Sergey Los

## Slide 3: MIP Measurements
Data Used: data/Skim/CdTe/Run1268-1310_RECO_NoRinging.root
Sensor was located behind the telescope inside the UCSC frame (roughly ~2m from center of telescope)
Track position computed at z = 200mm from center of telescope
Cuts: Photek amplitude > 0.1V and < 0.25V
CdTe channel passes “Ringing filter”
Track Position [X,Y] for those cuts

## Slide 4: MIP Measurements
Track position computed at z = 200mm from center of telescope
Cuts: Photek amplitude > 0.1V and < 0.25V
CdTe channel passes “Ringing filter”
Add CdTe amplitude > 0.12 mV cut
Sensor outline
(1cm x 1cm)
Note that the beam was not super well aligned with the sensor, and therefore we did not capture the full width of the sensor. Instead we only captured about 8mm in X, and 4.5mm in Y. The beam drops off below around Y=6mm, the pixel telescope coverage ends below X~11mm

## Slide 5: MIP Measurements
Track position computed at z = 200mm from center of telescope
Cuts: Photek amplitude > 0.1V and < 0.25V
CdTe channel passes “Ringing filter”
Require track position within fiducial area of the sensor:
X: (11mm, 18.5mm)
Y: (6.5mm, 11.3mm)

## Slide 6: MIP Measurements
MIP signal charge
I’m still concerned whether the baseline has been fitted well. If not then the charge measurement has a pedestal bias which will affect our interpretation of the measured value and the comparison with showers
Integrate a window around the peak
[-25 bins, +50 bins]
Peak and falling edge fits landau well, but low edge does not. Low edge may possibly have background?
Integrate the full sampling range (1024 bins)
Shape becomes gaussian too for some reason

## Slide 7: MIP Measurements
MIP signal amplitude
Note that the beam was not super well aligned with the sensor, and therefore we did not capture the full width of the sensor. Instead we only captured about 8mm in X, and 4.5mm in Y
MIP signal amplitude is stable vs particle position

## Slide 8: MIP Efficiency
Efficiency above 99%
Use amplitude > 0.04 mV to define numerator

## Slide 9: MIP Timing
Use pulses with amplitude > 0.03 mV
There’s some structure vs position. Need to check if it’s consistent with what we saw previously

## Slide 10: MIP Timing
Time resolution is about 450ps

## Slide 11: CdTe
Absorber | 5-7mm Air gap | Silicon Pad | 5-7mm Air gap | CdTe
Choose region:
X: 12 – 19 mm
Y: 16 – 23 mm

## Slide 12: CdTe
Absorber | 5-7mm Air gap | Silicon Pad | 5-7mm Air gap | CdTe
2X0 Lead

## Slide 13: CdTe
Absorber | 5-7mm Air gap | Silicon Pad | 5-7mm Air gap | CdTe
4X0 Lead

## Slide 14: CdTe
Absorber | 5-7mm Air gap | Silicon Pad | 5-7mm Air gap | CdTe
6X0 Lead

## Slide 15: CdTe
Absorber | 5-7mm Air gap | Silicon Pad | 5-7mm Air gap | CdTe
4.3X0 Tungsten + 6X0 Lead = 10.3 X0

## Slide 16: Silicon
Absorber | 5-7mm Air gap | Silicon Pad | 5-7mm Air gap | CdTe
Choose region:
X: 11.5 – 15 mm
Y: 13 – 15 mm
Silicon and CdTe sensor were offset with respect to each other.
Clearly we didn’t sample the center of the shower for the silicon case

## Slide 17: Silicon
Absorber | 5-7mm Air gap | Silicon Pad | 5-7mm Air gap | CdTe
4X0 Lead

## Slide 18: Silicon
Absorber | 5-7mm Air gap | Silicon Pad | 5-7mm Air gap | CdTe
6X0 Lead

## Slide 19: Silicon
Absorber | 5-7mm Air gap | Silicon Pad | 5-7mm Air gap | CdTe
4.3X0 Tungsten + 6X0 Lead = 10.3 X0

## Slide 20
Absorber | 5-7mm Air gap | CdTe | 5-7mm Air gap | Silicon Pad

## Slide 21: CdTe
Absorber | 5-7mm Air gap | Silicon Pad | 5-7mm Air gap | CdTe
Choose region:
X: 14 – 20 mm
Y: 24 – 26 mm
Alignment is off-center in Y
But it’s maybe enough to get some of the center region

## Slide 22: CdTe
Absorber | 5-7mm Air gap | CdTe | 5-7mm Air gap | Silicon Pad
4X0 Lead

## Slide 23: CdTe
Absorber | 5-7mm Air gap | CdTe | 5-7mm Air gap | Silicon Pad
6X0 Lead

## Slide 24: CdTe
Absorber | 5-7mm Air gap | CdTe | 5-7mm Air gap | Silicon Pad
4.3X0 Tungsten + 6X0 Lead = 10.3 X0

## Slide 25: Silicon
Absorber | 5-7mm Air gap | Silicon Pad | 5-7mm Air gap | CdTe
Choose region:
X: 10.5 – 13 mm
Y: 24 – 26 mm
Alignment is off-center in X and Y, but it’s maybe enough to get some of the center region

## Slide 26: Silicon
Absorber | 5-7mm Air gap | CdTe | 5-7mm Air gap | Silicon Pad
4X0 Lead

## Slide 27: Silicon
Absorber | 5-7mm Air gap | CdTe | 5-7mm Air gap | Silicon Pad
6X0 Lead

## Slide 28: Silicon
Absorber | 5-7mm Air gap | CdTe | 5-7mm Air gap | Silicon Pad
4.3X0 Tungsten + 6X0 Lead = 10.3 X0

## Slide 29: Shower Signal Comparison
Absorber | CdTe Mean Charge | Silicon Mean Charge | Mean Charge Ratio
2X0 Lead | 127 fC | No data | 
4X0 Lead | 337 fC | 119 fC | 2.8
6X0 Lead | 252 fC | 138 fC | 1.8
4.3W + 6X0Lead | 133 fC | 50 fC | 2.7
Absorber | CdTe Mean Charge | Silicon Mean Charge | Mean Charge Ratio
4X0 Lead | 312 fC | 47 fC | 6.6
6X0 Lead | 316 fC | 41 fC | 7.7
4.3W + 6X0Lead | 288 fC | 36 fC | 8.0
Absorber | 5-7mm Air gap | CdTe | 5-7mm Air gap | Silicon Pad
Absorber | 5-7mm Air gap | CdTe | 5-7mm Air gap | Silicon Pad

## Slide 30: Shower Signal Comparison
Absorber | CdTe Mean Charge | Silicon Mean Charge | Mean Charge Ratio
4X0 Lead | 312 fC | 119 fC | 2.6
6X0 Lead | 316 fC | 138 fC | 2.3
4.3W + 6X0Lead | 288 fC | 50 fC | 5.8
Absorber | CdTe Mean Charge | Silicon Mean Charge | Mean Charge Ratio
4X0 Lead | 337 fC | 47 fC | 7.2
6X0 Lead | 252 fC | 41 fC | 6.1
4.3W + 6X0Lead | 133 fC | 36 fC | 3.7
Sensor Located after first air gap
From previous slide, clearly the air gaps make a big difference
Let’s compare the signals under the same air gap distance
Sensor Located after 2nd air gap

## Slide 31: Shower Signal Comparison
Absorber | CdTe Mean Charge | Silicon Mean Charge | Mean Charge Ratio
4X0 Lead | 312 fC | 119 fC | 2.6
6X0 Lead | 316 fC | 138 fC | 2.3
4.3W + 6X0Lead | 288 fC | 50 fC | 5.8
Absorber | CdTe Mean Charge | Silicon Mean Charge | Mean Charge Ratio
4X0 Lead | 337 fC | 47 fC | 7.2
6X0 Lead | 252 fC | 41 fC | 6.1
4.3W + 6X0Lead | 133 fC | 36 fC | 3.7
Sensor Located after first air gap
At 10X0, CdTe seems to see a bigger signal that disappears if we place it after the silicon sensor + copper box + air gaps
  Maybe the x-ray photons get absorbed in the copper / air?
Sensor Located after 2nd air gap

## Slide 32: Shower Signal Comparison
Absorber | CdTe Mean Charge | Silicon Mean Charge | Mean Charge Ratio
4X0 Lead | 312 fC | 119 fC | 2.6
6X0 Lead | 316 fC | 138 fC | 2.3
4.3W + 6X0Lead | 288 fC | 50 fC | 5.8
Absorber | CdTe Mean Charge | Silicon Mean Charge | Mean Charge Ratio
4X0 Lead | 337 fC | 47 fC | 7.2
6X0 Lead | 252 fC | 41 fC | 6.1
4.3W + 6X0Lead | 133 fC | 36 fC | 3.7
Sensor Located after first air gap
These ratios are too small compared to expectation ( even if there are no x-ray photons, ratio should be ~5)
Need to understand if it’s because silicon charge is biased by imperfect pedestal subtraction. It’s better if we can amplify the silicon more so that it suffers less from pedestal
Sensor Located after 2nd air gap

## Slide 33: Shower Signal Comparison
Absorber | CdTe Mean Charge | Silicon Mean Charge | Mean Charge Ratio
4X0 Lead | 312 fC | 119 fC | 2.6
6X0 Lead | 316 fC | 138 fC | 2.3
4.3W + 6X0Lead | 288 fC | 50 fC | 5.8
Absorber | CdTe Mean Charge | Silicon Mean Charge | Mean Charge Ratio
4X0 Lead | 337 fC | 47 fC | 7.2
6X0 Lead | 252 fC | 41 fC | 6.1
4.3W + 6X0Lead | 133 fC | 36 fC | 3.7
Sensor Located after first air gap
Also the ratio changes a lot between the first air gap and 2nd air gap.
Is it due to spread of shower in transverse direction? The silicon pad is smaller in area compared to CdTe.
Sensor Located after 2nd air gap

## Slide 34: Summary
ss

## Slide 35: Backups